# Vagal Nerve Stimulator (VNS)

Last updated: September 18, 2024

Comparison of Neuromodulations (RNS, DBS, VNS) – see p. E11 >>

Transcutaneous VNS	3

Surgical VNS	3

Manufacturers	4

USA	4

Cyberonics → LivaNova	4

MicroTransponder Inc.	4

China	4

PINS	4

Indications	5

Epilepsy	5

Depression	6

Rehabilitation after ischaemic stroke	6

Investigational	6

Heart failure	6

Contraindications	7

Mechanism of action	7

Cardioprotection	7

Seizures	7

New Paradigms	9

Microburst VNS therapy	9

leads - LivaNova	12

VNS Therapy (model 302)	14

PerenniaDURA (model 303)	14

PerenniaFLEX (model 304)	14

Hardware (component compatibility)	15

leads - Vivistim	15

Implantable pulse generators (IPG) - PINS	15

Implantable pulse generators (IPG) - Vivistim	16

Implantable pulse generators (IPG) - LivaNova	16

Battery capacity	18

Legacy (non-available)	19

NCP (models 101, 101)	19

Pulse (model 102)	19

Pulse Duo (model 102R) -  dual pin	20

Aspire HC (model 105)	21

Legacy (still available)	22

Demipulse (model 103)	22

Demipulse Duo (model 104) - dual pin	23

Heart rate sensing (AutoStim)	23

Aspire SR (model 106)	23

Sentiva (model 1000)	24

Sentiva Duo (model 1000-D) – dual pin	25

Depression	26

Symmetry (model 8103)	26

Surgical accessories (LivaNova)	26

Magnet (LivaNova)	27

Programmer (LivaNova)	28

Preoperative work-up	28

New Implantation	28

Battery replacements	29

Battery indicators	29

How to handle depleting battery - Recommendations for neurosurgeon	29

Lead Revision / Broken lead	30

Explantation of entire system	31

Surgical Technique – new VNS implantation	31

General Schema	32

Anatomy	33

Position	34

Incision	35

Dissection	35

Nerve Exposure	35

IPG pocket in the chest wall	36

Lead Tunneling	37

Wrapping Lead	38

Pulse Generators	43

Testing	44

Intraoperative programming	46

Closure	46

Postoperatively	46

Patient Instructions	47

MRI	48

Group A - 103, 105, 106, 1000 implanted higher than 4th rib (armpit level)	48

Group B – 102, 104, also Group A implanted below 4th rib	49

Cut / Broken Lead	51

Device preparation for MRI scan	54

Electrosurgery	54

Programming	54

Stimulation waveform	54

Strength–duration time constant (chronaxie)	55

Canonical biophysical laws	55

Intraoperative	55

Magnet Mode	56

AutoStim – Tachycardia Detection	56

Bradycardia Detection – “SUDEP monitor”	57

Impedance	57

Clinic - Programming Phase I	58

Output Current	58

Pulse Width	59

Frequency	59

Clinic - Programming Phase II	62

Duty Cycle	62

Side Effects	63

Complications	64

Drug interactions	64

Outcomes - Depression	65

Outcomes - Epilepsy	65

Different epilepsy types	70

Generalized Epilepsy	71

Hippocampal Epilepsy	71

Tumor-associated epilepsy	72

Cost-utility analysis	72

Transcutaneous VNS

Surgical VNS

VNS - palliative continuously intermittent electrical stimulation with chest-wall-implanted pacemaker-like device.

Pending Reading

- Morris, G. L., Gloss, D., Buchhalter, J., Mack, K. J., Nickels, K., & Harden, C. (2013). Evidence-based guideline update: vagus nerve stimulation for the treatment of epilepsy: report of the Guideline Development Subcommittee of the American Academy of Neurology. Neurology, 81(16), 1453-1459.

- J. Wheless, A. Gienapp, P. Ryvlin. Vagus nerve stimulation (VNS) update. Epilepsy Behav, 88 (2018), pp. 2-10.

- P. Ryvlin, E.L. So, C.M. Gordon, D.C. Hesdorffer, M.R. Sperling, O Devinsky, et al. Long-term surveillance of SUDEP in drug-resistant epilepsy patients treated with VNS therapy. Epilepsia, 59 (2018), pp. 562-572

- Brodie MJ. Curr Neurol Neurosci 2016.

Manufacturers

USA

Cyberonics → LivaNova

Cyberonics >>  → LivaNova >> also >>

Full Manual >>

- first models were made by Cyberonics company >>

- now made by LivaNova company >>

LivaNova October 26, 2023 data breach

LivaNova discovered incident that resulted in a disruption to portions of their IT systems - unauthorized party obtained certain personal information: name, contact information (e.g., phone number, email and postal address), date of birth, medical information (e.g., treatment, condition, diagnosis, prescription, physician, medical record number and device serial number), and health insurance information.

MicroTransponder Inc.

https://www.vivistim.com

China

PINS

- Chinese company Beijing Pins Medical Co., Ltd also manufactures VNS systems that have Bluetooth connectivity and patients have remote to communicate with their device; not yet approved in Europe or USA >>

Indications

Epilepsy

- focal-onset epilepsy (even with anatomically defined single foci); FDA approval (1997): “adjunctive therapy for reducing the frequency of seizures in patients > 4 years with partial-onset seizures refractory to antiepileptic medications”

- some insurances may deny if not for "partial epilepsy" (ICD-10 codes: G40.211 / G40.219 / G40.011 / G40.019 / G40.111 / G40.119)

- Medicare / CMS National Coverage Determination (NCD) 160.18 for VNS states: "Effective for services performed on or after July 1, 1999, VNS is reasonable and necessary for patients with medically refractory partial onset seizures for whom surgery is not recommended or for whom surgery has failed."

- reports of implantation in 1 year-old patient (word of caution – as child grows, the lead may start pulling on vagus nerve – cases of bradycardia that resolved after lead replacement).

- mounting evidence suggests effectiveness in generalized epilepsy (VNS is one of preferred treatments for Lennox-Gastaut syndrome)

N.B. on label is only for partial epilepsy! (off label can use for generalized epilepsy – VNS works very well!!!)

- VNS has been claimed to effectively control atonic seizures as an alternative to corpus callosotomy (meta-analysis of 31 studies involving 533 children showed that callosotomy was more effective than VNS at a cost of greater adverse events, including twice as many reoperations and a 14% rate of symptomatic disconnection syndrome).

- stimulation of either vagus nerve is effective, but left nerve is always chosen (theoretically, less likely to cause cardiac effects), usually left cervical vagus nerve.

- if left nerve is unavailable, there are case reports of implantation on the right side without side effects.

  - dog studies - stimulation of the right vagus nerve caused more cardiac slowing then the left.

Rutecki P. Anatomical, physiological, and theoretical basis for the antiepileptic effect of vagus nerve stimulation. Epilepsia 1990;31:S1–6

N.B. VNS can be combined with RNS, DBS.

VNS, RNS, and DBS are all palliative and comparable in efficacy, both in pivotal trials and over longer-term trials. VNS is often a first choice as it is extracranial.

Depression

FDA approved (2007) VNS as adjunctive long-term treatment of chronic / recurrent depression in patients > 18 years with a major depressive episode not adequately relieved by ≥ 4 antidepressant treatments.

N.B. improvement in mood may be a huge additional benefit for epilepsy patients!

- depression prevalence is 13–37% of patients with epilepsy (number is higher in uncontrolled epilepsy).

- antidepressive effect is independent of seizure benefit.

Symmetry (model 8103)  (LivaNova) technical details – see below >>

Rehabilitation after ischaemic stroke

Read articles

- Dawson J, et al. Vagus nerve stimulation paired with rehabilitation for upper limb motor function after ischaemic stroke (VNS-REHAB): a randomised, blinded, pivotal, device trial. Lancet, 2021:397(10284), 1545.

- MicroTransponder data on file.

- PMA P210007

- Francisco GE, et al. Vagus nerve stimulation paired with upper-limb rehabilitation after stroke: two- and three-year follow-up from the pilot study. Arch Phys Med Rehabil, 2023:104(8), 1180.

- Dawson J, et al. Vagus nerve stimulation paired with rehabilitation for upper limb motor impairment and function after chronic ischemic stroke: subgroup analysis of the randomized, blinded, pivotal, VNS-REHAB device trial. Neurorehabilitation and Neural Repair, 2023:37(6), 367.

- Meyers EC, et al. Vagus nerve stimulation enhances stable plasticity and generalization of stroke recovery. Stroke. 2018;49(3):710.

Vivistim® Paired VNS™ System (MicroTransponder Inc) is intended to be used e during rehabilitation therapy to improve motor function in chronic ischemic stroke patients with moderate to severe arm impairment.

Vivistim technical details – see below >>

Investigational

Heart failure

- Autonomic Regulation Therapy (ART) delivered through VNS. ART is being studied in a clinical trial called ANTHEM-HFrEF.

Contraindications

- Absence of vagus nerve, history of cervical vagotomy.

- Absence of epilepsy / depression

- Not surgical candidate

- Vocal cord paralysis – relative contraindication directing to use vagus nerve on the opposite (right) side.

- Distorted neck anatomy (history of infection, radiotherapy, radical neck dissection) – relative contraindication directing to use vagus nerve on the opposite (right) side.

Additional precautions are necessary in pregnancy, chronic pulmonary diseases such as asthma an COPD, obstructive sleep apnea (OSA), cardiac bradyarrhythmias, active peptic ulcer disease, and insulin dependent diabetes mellitus.

- patient with OSA - use night-time lower stim parameters (with Sentiva).

Mechanism of action

- 20-40% of vagus fibers are efferent – cause side effects;

nucleus ambiguus → skeletal laryngeal musculature, bradycardia or heart block.

dorsal nucleus of the vagus → GI side effects and anorexia.

- VNS works via GABAergic mechanism.

Cardioprotection

- VNS stimulation reduces T-wave alternans via anti-adrenergic effects, improved baroreflex sensitivity, anti-inflammatory / anti-apoptosis effects.

Epilepsy ≈ 3-fold increased risk of lethal cardiac arrhythmias!

Seizures

- 80% of vagal fibers are afferent → synapse mainly onto the nucleus of the solitary tract → locus coeruleus and dorsal raphe → ↑CNS noradrenergic tone (norepinephrine is anti-epileptic)

Locus coeruleus - key central anti-convulsive target of VNS; VNS likely works through thalamus

- project to many structures in brain (incl. hippocampus, amygdala, thalamus).

- changes are seen in thalamic blood flow that correlate with efficacy.

- decline in interictal spikes can be seen on serial EEGs with chronic VNS.

- SPECT shows normalization of GABAA receptor density in patients with good response to 1 year of VNS stimulation.

- in numerous animal studies, VNS had both anticonvulsive and proconvulsive effects.

- it seems that, as in human studies, focal seizure might be influenced in 30% of cases, and that in cases of generalized epilepsy, VNS can show improvement or at the opposite, no effect and sometime worsening.

- animal experiments cannot so far answer the following question: does efficacy improve after chronic stimulation?

Pathways affected by VNS:

2. Koo B. J Clin Neurophysiol. 2001; 434-441. 3. Roosevelt et al. Brain Res. 2006; 124-132. 4. Ben-Menachem et al. Epilepsy Res. 1995;221-227. 5. Vonck et al. Seizure. 2008.05.001.

VNS and fMRI - predict response to VNS based on cortico-thalamic connectivity (article by George Ibrahim from Sick Kids).

New Paradigms

Microburst VNS therapy

- high frequency bursts of electrical stimulation, called microbursts.

Animal study - rCBF

Martlé V et al. Regional brain perfusion changes during standard and microburst vagus nerve stimulation in dogs. Epilepsy Res. 2014 May;108(4):616-22.

VNS Therapy System was implanted in ten Beagle dogs. μ-SPECT was performed after sham, standard and microburst VNS in a randomized, cross-over study. Nineteen volumes of interest (VOIs) were semi-quantitatively analysed and perfusion indices (PIs) were calculated. Furthermore, a rostro-caudal gradient (R-C), an asymmetry index (AI) and a cortical-subcortical index (Co-SCo) were determined. The SPECT results after standard and microburst VNS were compared pairwise with sham stimulation.

RESULTS:

Acute standard VNS did not cause significant rCBF alterations. Acute microburst VNS caused a significant hypoperfusion in the left frontal lobe (P=0.023) and in the right parietal lobe (P=0.035). Both stimulation paradigms did not cause changes in R-C, AI nor Co-SCo.

CONCLUSIONS:

Microburst VNS is more potent than standard VNS to modulate the rCBF in the dog. Our results promote further research towards the antiepileptic effect of microburst VNS in dogs and humans.

Animal study - transmitter

Martlé V et al. The effect of vagus nerve stimulation on CSF monoamines and the PTZ seizure threshold in dogs. Brain Stimul. 2015 Jan-Feb;8(1):1-6.

OBJECTIVE:

We aimed to investigate the effect of two rapid cycling VNS paradigms on CSF monoamine levels and the seizure threshold in the canine pentylenetetrazole (PTZ) model.

METHODS:

Eight Beagle dogs, implanted with a VNS Therapy(®) System, participated in a cross-over study. Levels of serotonin (5HT), norepinephrine (NE) and dopamine (DA) were quantified in the CSF after 1 h of sham, standard and microburst VNS with a wash-out period of 1 month. One week after the CSF experiment, the PTZ seizure threshold was determined after the same stimulation paradigm. As a positive control, the PTZ seizure threshold was determined after a single oral dose of phenobarbital.

RESULTS:

Rapid cycling standard and microburst VNS caused a significant increase of NE levels in the CSF (P = 0.03 and P = 0.02 respectively). No significant changes in 5HT or DA levels were detected. Rapid cycling standard and microburst VNS did not cause significant changes in the PTZ seizure threshold compared to sham.

CONCLUSIONS:

VNS induces an increase of NE in the canine brain, which supports previous findings indicating that VNS influences the locus coeruleus-NE (LC/NE) system. Importantly, this study demonstrates that this increase in NE is measurable in the CSF. One hour of VNS did not affect seizure threshold in the canine PTZ model. Therefore, the role of NE in the antiepileptic effect of VNS in dogs remains to be elucidated.

Animal study – heart rate

Martlé V et al. Evaluation of heart rate variability in dogs during standard and microburst vagus nerve stimulation: a pilot study. Vet J. 2014 Dec;202(3):651-3.

The present study investigated heart rate variability (HRV) in healthy Beagle dogs treated with 1 h of sham, standard or microburst left-sided VNS in a crossover design. No significant differences were found between the stimulation paradigms for any of the cardiac parameters. Short-term left-sided VNS, including a novel bursting pattern (microburst VNS), had no statistically significant effect on HRV in ambulatory healthy dogs.

Humans - Feasibility Study

ClinicalTrials.gov Identifier: NCT03446664

| Principal Investigator: | Selim Benbadis, MD | University of South Florida Health |
| --- | --- | --- |

| Contact: Amy Keith | 281-228-7495 | Amy.keith@livanova.com |  |
| --- | --- | --- | --- |
| Contact: Jeffrey Way | 281-228-7394 | Jeffrey.way@livanova.com |  |

| Estimated Enrollment : | 40 participants |
| --- | --- |
| Intervention Model: | Single Group Assignment |
| Intervention Model Description: | Two cohorts of subjects with refractory epilepsy; (1) subjects with primary generalized tonic-clonic seizures and (2) subjects with partial onset seizures including complex partial seizures with or without secondary generalization. |
| Masking: | None (Open Label) |
| Primary Purpose: | Treatment |
| Actual Study Start Date : | February 27, 2018 |
| Estimated Primary Completion Date : | April 2021 |
| Estimated Study Completion Date : | June 2021 |

Primary Outcome Measures

Efficacy Primary Endpoint: Percent change from baseline in seizure frequency [Time Frame: Up to 12 months study visit]

the change in the seizure frequency per month compared to baseline will be evaluated for each subject at follow-up visits month 6 and 12.

Safety Primary Endpoint: Occurrence of stimulation related Adverse Events [Time Frame: Up to 12 months study visit]

Assess stimulation/device related adverse events at follow-up visits month 6 and 12.

Secondary Outcome Measures

Change from baseline in seizure frequency per month based on seizure diary provided by the sponsor [ Time Frame: Up to 12 months study visit ]

Change from baseline in seizure severity [ Time Frame: Up to 12 months study visit ] - as measured by the Seizure Severity Questionnaire (SSQ) scale (Cramer, 2002).

Change from baseline in quality of life [ Time Frame: Up to 12 months study visit ] - as measured by the QOLIE-31-P for adults 18 years and older (Cramer et al.; 1998) and QOLIE-AD-48 for adolescents 12 to 17 years (Cramer et al.; 1999).

Change from baseline in antiepileptic drug (AED) load [ Time Frame: Up to 12 months study visit ]

Estimated as the sum of the prescribed daily dose (PDD)/defined daily dose (DDD) ratios for each AED included in the treatment regimen (Deckers et al., 1997), where DDD (WHO ATC/DDD index) corresponds to the assumed average therapeutic daily dose of a drug used for its main indication.

Suicidality as measured by the Columbia Suicide Severity Rating Scale (C-SSRS) [ Time Frame: Up to 12 months study visit ]

All adverse events [ Time Frame: Up to 12 months visit ]

Eligibility Criteria

| Ages Eligible for Study: | 12 Years and older (Child, Adult, Older Adult) |
| --- | --- |
| Sexes Eligible for Study: | All |
| Accepts Healthy Volunteers: | No |

Inclusion Criteria:

- Clinical diagnosis of medically refractory epilepsy with primary generalized tonic-clonic seizures (limited to 20 subjects) or partial onset seizures including complex partial seizures with or without secondary generalization (limited to 20 subjects).

- Must be on adjunctive antiepileptic medications.

- Willing and capable to undergo multiple evaluations with functional magnetic resonance imaging (fMRI), electroencephalogram (EEG) and electrocardiogram (ECG).

- 4(A) For subjects with partial onset seizures: An average of ≥ 3 countable seizures per month based on seizure diary during the 3 month baseline period and no seizure-free interval greater than 30 days during those 3 months.

4(B) For subjects with PGTCs: Have at least ≥ 3 countable seizures during the 3 month baseline period. Note: Each seizure within a cluster may be counted as separate seizures.

- 12 years of age or older.

- Subject is a male or non-pregnant female adequately protected from conception. Females of childbearing potential must use an acceptable method of birth control.

- Provide written informed consent-assent/Health Insurance Portability and Accountability Act (HIPAA) authorization and self-reported measures with minimal assistance as determined by the investigator.

Exclusion Criteria:

- Currently using, or are expected to use, short-wave diathermy, microwave diathermy, or therapeutic ultrasound diathermy.

- A VNS Therapy System implant would (in the investigator's judgment) pose an unacceptable surgical or medical risk for the subject.

- A planned procedure that is contraindicated for VNS therapy.

- History of implantation of the VNS Therapy System.

- Currently receiving treatment from an active implantable medical device.

- Presence of contraindications to MRI per the MRI subject screening record.

- Known clinically meaningful cardiovascular arrhythmias currently being managed by devices or treatments that interfere with normal intrinsic heart rate responses (e.g., pacemaker dependency, implantable defibrillator, beta adrenergic blocker medications).

- History of chronotropic incompetence (commonly seen in subjects with sustained bradycardia [heart rate < 50 bpm]).

- Cognitive or psychiatric deficit that in the investigator's judgment would interfere with the subject's ability to accurately complete study assessments.

- History of status epilepticus within 1 year of study enrollment.

- Dependent on alcohol or narcotic drugs as defined by DSM IV-TR within the past 2 years, based on history. Tests for drug or alcohol use will not be administered.

- Currently being treated with prescribed medication that contains cannabis or cannabis related substance.

- Any history of psychogenic non-epileptic seizures.

- Currently participating in another clinical study without LivaNova written approval.

leads - LivaNova

Only dual-pin lead is model 300 – no longer manufactured.

Lead characteristics: insulation – silicone; resistance (pin-to-electrode) – 120 to 250 Ohm; conductor material – platinum iridium.

Each lead is available in two sizes (2.0- and 3.0 mm electrode inner diameter) to ensure optimal electrode fit on different size nerves.

The lead’s lifespan is undetermined by manufacturer. A lead would require replacement if a lead fracture were suspected, accompanied by increased symptoms (e.g., seizure frequency).

List price for 304 – 8,068 USD (2020 March), 6,744.75 USD (June 2024)

VNS Therapy (model 302)

PerenniaDURA (model 303)

PerenniaFLEX (model 304)

Hardware (component compatibility)

leads - Vivistim

Implantable pulse generators (IPG) - PINS

Implantable pulse generators (IPG) - Vivistim

Implantable pulse generators (IPG) - LivaNova

Currently available models are in two footprint sizes:

smaller – 103, 1000

larger – 105, 106

- hermetically sealed titanium case with polyurethane header and stainless steel lead connector blocks.

Battery capacity

103, 104, 1000 - 1 Ah

102, 105, 106 - 1.7 Ah

- self-discharge reduces the capacity by < 1% per year.

- battery lasts for 10,000 hours of stimulation.

- battery performs equally at 100% and at 1% (but it is not always true clinically – sometimes when battery is < 10% patients start having worsening seizures – unexplained phenomenon that has no basis) further see >>

- no harm leaving dead battery in - will not leak.

Legacy (non-available)

NCP (models 101, 101)

Source of picture: LivaNova >>

Pulse (model 102)

Source of picture: LivaNova >>

Pulse Duo (model 102R) -  dual pin

Source of picture: LivaNova >>

Aspire HC (model 105)

Source of picture: LivaNova >>

Legacy (still available)

(limited for existing VNS battery replacements)

List price – 27,090 USD (2020 March)

Demipulse (model 103)

Source of picture: LivaNova >>

Demipulse Duo (model 104) - dual pin

List price – 27,090 USD (2020 March)

Source of picture: LivaNova >>

Heart rate sensing (AutoStim)

about AutoStim – see below >>

Aspire SR (model 106)

List price – 31,557 USD (2020 March), 26,094.75 USD (June 2024)

Source of picture: LivaNova >>

Sentiva (model 1000)

List price – 35,271 USD (2020 March), 29,358.75 USD (June 2024)

Read 3rd page of brochure!!!! >>

December 20, 2019 | The FDA announced a recall of the LivaNova VNS Therapy SenTiva Generator due to an unintended reset error that causes the system to stop delivering VNS therapy.

https://www.fda.gov/medical-devices/medical-device-recalls/livanova-recalls-vns-therapy-sentiva-generator-due-reset-error

2020/03/03 LivaNova update:

- reset occurred in 0.5% of devices and only within first 60 days of enabling therapy – all outputs spontaneously reset to 0 mA.

- currently, all new devices are screened and no longer have this problem after firmware update (“LivaNova is currently distributing Model 1000 devices that have passed a LivaNova internal error screen. The error screen is intended to detect devices susceptible to unintended device disablement“).

Patients can check their Sentiva device status by entering SN (serial number):

https://www.vnssentivareset.com

- concerned patients also can use magnet regularly to see if device is working.

Source of picture: LivaNova >>

- new features:

- Low Heart Rate detection*

- Prone Position detection*

- Guided Programming + Scheduled Programming Titration schedule - can be programmed during one office visit and delivered while the patient lives their life (e.g. increase stim current every 2 weeks at the specific time of the day until therapeutic target is reached).

- Day & Night Programming - two independent sets of parameters can be customized and delivered based on each patient’s needs.

- Events & Trends - quickly view data at follow-up visits on events that may be associated with seizures.

*does nothing therapeutically (does not trigger stim); activated only if device “thinks” it was a seizure (patient swiped a magnet or AutoStim went off) – bradycardia or prone position may be risk factors for SUDEP – Sentiva VNS acts as “SUDEP monitor”

Sentiva Duo (model 1000-D) – dual pin

- for cases with legacy dual-pin leads still in place.

List price – 29,358.75 USD (June 2024)

Source of picture: LivaNova >>

Depression

Symmetry (model 8103)

- specifically designed for depression.

- footprint same as 103, 1000.

- does not detect heart rate.

Source of picture: LivaNova >>

Surgical accessories (LivaNova)

List price for tunneler (#402) – 506 USD (2020 March)

Source of picture: LivaNova >>

List price for accessory kit (#502) – 721 USD (2020 March)

- has a spare screwdriver (one screw driver comes in the package with IPG) and 4000 Ohms test resistor.

Source of picture: LivaNova >>

Magnet (LivaNova)

With each IPG implantation, the patient gets two identical VNS Therapy magnets (#220), each providing a minimum of 50 gauss at 1 inch:

Source of picture: LivaNova >>

Programmer (LivaNova)

- the treating health care professional gets from LivaNova company a clinician programmer with a wireless wand:

Source of picture: LivaNova >>

- future models should have Bluetooth technology - device access to the patient on the personal cell phone and remote data access to the health provider.

Preoperative work-up

New Implantation

- confirm a well-established diagnosis of pharmacoresistant epilepsy - review the brain MRI, ideally conducted on a 3 Tesla scanner with an epilepsy protocol, and carefully examine the scalp EEG report.

- inquire about history of obstructive sleep apnea, as well as their occupation and hobbies that may require a clear voice.

- assess history of any prior neck surgeries and inquire about heart problems.

- auscultation for carotid bruits.

- ECG – look for AV node or bundle branch blocks – high risk of cardiac side effects during VNS stimulation.

- ensure effective patient counseling regarding the anticipated side effects of VNS, as outlined below.

Battery replacements

- if not sure if VNS has efficacy, turn down output to 1.0 mA and watch seizure response (better strategy than letting VNS to expire and then watch – liability issue if patient goes into status).

Battery indicators

- OK - normal operating range and no special attention is required

75-100%

50-75%

25-50%

11-25%

- Intensified Follow-up Indicator (IFI) - battery has depleted to a level where more frequent clinical monitoring is recommended:

5-11% (model 105, 106) or 8-18% (model 103, 104)

- Near End of Service (N EOS) - generator should be replaced as soon as possible but battery still maintains 3.3 V.

- End of Service (EOS) - generator is no longer supplying stimulation and immediate replacement is recommended.

- battery starts gradual drop in voltage and generator will eventually lose the ability to communicate with the clinician programmer (pre-operative impedance check will not be possible).

How to handle depleting battery - Recommendations for neurosurgeon

- see patient in clinic when battery is IFI.

- educate patient:

- ideally, battery replacement should occur at NEOS stage.

- every surgery has risks (incl. the risk of infection leading to VNS explantation).

- too early battery replacement:

- unnecessarily exposes to surgical risks

- increases number of surgeries in patient’s lifetime

- battery depletion can occur between office visits - recommend patients to perform a daily magnet activation to check stimulation (if stimulation is not felt, instruct the patient to consult with the physician to perform diagnostics testing).

- calculate time to NEOS:

- LivaNova engineers can calculate BLC (battery life calculated) – you need to provide current interrogation data:

- call 866-882-8804

- email clinicaltechnicalservices@livanova.com

- Surgeon can use table to calculate intervals from IFI to NEOS; e.g. for model 106 >>

- schedule battery replacement surgery near predicted NEOS time (any troubles of calculating time to NEOS – schedule frequent visits with either you or referring neurologist).

- use this waiting time to review MRIs and EEGs – it is not uncommon to encounter patients with VNS and lesional MRI:

- VNS was placed inappropriately [previous provider did not recognize curable etiology of epilepsy or modern surgery was not available].

- patient refused resective surgery in the past whereas modern approaches [LITT, RNS, DBS] can be more appealing to the patient today due to minimally invasive approach.

- avoid inertia when the original VNS was implanted but no efficacy was ever observed / documented; however, subsequent replacements are done “because patient has VNS and battery is low”.

- surgeon has to personally interrogate system preoperatively for impedance and battery level, also to make sure the implanted lead and IPG models are known preoperatively as the older electrode designs (two pin) will require legacy IPG models for compatibility. If information is missing and IPG battery is completely depleted, perform X-ray to identify the system.

- ensure that the informed consent includes VNS lead revision if indicated.

Lead Revision / Broken lead

Impedances

- either too high (> 10,000 Ohm) or too low (esp. < 600 Ohm).

- look for historic impedances – sudden changes (esp. significant decrease in impedance) are strongly suggestive.

Clinical features

- increase in seizures.

- patient perception of feeling erratic, limited, no stimulation.

- short-circuit condition - painful feeling in the neck (that stops after turning VNS off).

Diagnosis

- X-ray: lead discontinuities can sometimes be identified on x-ray of the VNS system.

- ECG – if patient is not receiving therapy (e.g. due to high impedances), it needs to be treated as new system implantation: look for AV node or bundle branch blocks preop.

- experimental - skin electrodes on the neck + either evoked potential monitoring equipment or an oscilloscope can be used to analyze the stimulus waveform for verification of an electrical discontinuity; differentiated waveform with narrowed pulses or no waveform at all can confirm a discontinuity:

Surgery

- no need to explant entire old lead (to jeopardize nerve) as patient’s MRI conditionality will not change due to the presence of VNS system (i.e. still no MRI in exclusion zone – see below >>).

- start with low output stimulation postop (as with new system).

- transient hoarseness may happen due to vagus manipulation.

Explantation of entire system

- if VNS has no efficacy, VNS is still left in situ – new IPGs in the future may offer new stimulation protocols.

- if patient insists on explantation (e.g. claim no efficacy or needs MRI in exclusion zone), turn stim off for > 3 months to make sure VNS is indeed not helping.

- remove as much of the lead as possible if explanting a system (ideally, should be < 2 cm remaining to allow MRI in exclusion zone – see below >>); transect lead here:

Surgical Technique – new VNS implantation

General Schema

Source of picture: Livanova >>

Anatomy

Source of picture: Livanova >>

- nerve usually lies in posterior groove between carotid artery and internal jugular vein.

Position

- head is rested on a gel donut / towels (Dr. Holloway)

- Dr. Holloway places saline bag under neck to push cervical viscera anteriorly.

- head rotation:

- straight (Dr. Holloway – as any rotation distorts anatomy!!!)

- rotate head to the right (Dr. Tye) – moves chin away from surgical access but internal jugular vein tends to overlap vagus nerve.

- arm position depends where IPG is going to be placed:

- arms tucked at the sides – for classical IPG implantations just below clavicle.

- right arm tucked at the side while left arm abducted on the arm rest (Dr. Tye for female patients if planning IPG incision lateral to breast; Dr. Collins places IPDs along anterior axillary fold).

- in the case of a twiddler patient, consider implanting IPG in the back.

Incision

- at level of cricothyroid membrane (“as for C5-6 ACDF or lower if neck anatomy allows – goal to avoid cardiac nerves”).

- 3-4 cm length of incision should be centered where carotid pulse is palpable, generally, with some overlap over medial edge of sternocleidomastoid muscle.

- inject local anesthetic with epinephrine early – allows hemostatic action to begin while sterile draping the patient (minimize electrosurgical use to occasional bipolar cautery use during dissection; should not need to use monopolar Bovie cautery at all!)

Dissection

- after reaching platysma, may cut it transversely (parallel to skin incision) using scissors or may split platysma along fibers (perpendicular to skin incision - Dr. Collins).

- Dr. Collins retracts omohyoid inferiorly using vessel loop.

- dissect “as for 2-level ACDF” – down to spine along medial edge of carotid; then dissect laterally along anterior surface of carotid until finding groove between carotid and IJ (do not dissect beneath SCM and do not expose IJ – IJ is fragile + does not need to be exposed at all!)

- Dr. Collins leaves some adventitia (edges of carotid sheath) on carotid and IJ (to engage retractor blades).

Nerve Exposure

- using delicate surgical instruments, for example Gerald tissue forceps, right-angle dissector, and Metzenbaum scissors, expose left vagus nerve half-way between clavicle and mastoid process (below where superior and inferior cervical cardiac branches separate from vagus nerve - stimulation of either of these two branches may cause bradycardia and/or asystole).

- the main vagus nerve trunk is the largest nerve in the area and typically resides in the posterior groove between the carotid artery and the internal jugular vein;

- one may encounter a thin flat nerve - superior root of ansa cervicalis – situated anteriorly between IJ and carotid):

N.B. based on 50 cadaver study for vagus syntopy, in 68% of cases, left vagus is situated anterior to common carotid (on the right, in 64% of cases vagus is posterior to CCA); comparing both sides, there was no symmetry in this syntopy. No influence of ethnic or anthropometric characteristics was observed

- be careful to never crush the nerve in the forceps – lift / manipulate the nerve using epineurium.

- dissect nerve circumferentially.

- expose 3 cm of vagus nerve - facilitates placement of lead helices on nerve.

N.B. ensure perfect hemostasis – use Gelfoam with thrombin (do not use electrocautery on or near nerve!)

- stretching nerve or allowing it to dry may result in temporary swelling → constriction → vocal cord dysfunction.

- Dr. Collins skeletonizes nerve thoroughly – removes all epineurium to minimize impedances (to conserve VNS battery) – he aims towards < 1500 ohm; risk – disturbing vasa nervorum and inducing later fibrosis.

IPG pocket in the chest wall

- this step takes only 1-2 minutes, so it is advisable to postpone it until the vagus nerve is fully exposed.

- classically, IPG is implanted just below the clavicle.

- traditionally, the IPG is implanted in a subcutaneous pocket just below the clavicle in the left upper chest. Try to place the generator at rib 4 or above, so the patient can have the maximum flexibility for MRI post-operatively.

- to determine the exact IPG location, palpate anterior chest wall tissues – find where there is enough tissue bulk to cover IPG (to maximize cosmesis, to minimize risk of skin dehiscence).

- skin incision can be perpendicular to clavicle or parallel to clavicle; in either case, ensure that incision is two-fingerbreadths away from clavicle and two-fingerbreadths away from deltopectoral groove.

- deepen subcutaneous dissection until pectoralis fascia is reached. Then continue dissection (including atraumatic blunt finger dissection along tissue plane) parallel and above the fascia in the direction of planned IPG pocket; may also create a short tissue flap on the opposite direction – to allow skin mobilization for tension-less closure.

- to avoid IPG communication difficulties, the pocket depth should not be deeper than 1 inch beneath the skin; similarly, it is not recommended to implant the IPG below muscle.

Lead Tunneling

- may do this step before wrapping lead contacts around nerve – less risk of lead dislodgement during tunneling.

- place bullet-tip end on Tunneler and tunnel subcutaneously.

- never route lead through muscle; never suture lead to muscle.

- after bullet tip has passed from one incision to other, unscrew bullet and withdraw shaft from sleeve, leaving sleeve extended through both incisions.

- use protective plastic sheath:

  - insert lead connector(s) inside end of sleeve at neck incision (if still using a legacy dual-pin lead, second connector will form slight compression fit between first lead connector tubing and sleeve inside).

  - pull sleeve, along with lead connector(s), towards chest incision until lead connector(s) completely exit(s) chest incision.

Source of picture: Livanova >>

  - Dr. Collins prefer using two protective plastic sleeves - one inside another: tunnel the large sleeve and then pull the smaller sleeve (with VNS lead inside) through it - this method protects the lead from pulling trauma, plus, when placing lead, no need to fight with lead memory and torque.

Wrapping Lead

- choose appropriately sized lead (2.0 or 3.0 mm electrode inner diameter) - should fit snugly without constricting nerve; 2.0 mm lead should accommodate most nerves.

  - it is not possible to predict what size lead will be needed.

  - open lead from its package only when ready to implant (i.e. do not expose lead to dust or other similar particulates, because its silicone insulation can attract particulate matter).

  - do not soak lead in saline or similar solution before implanting it - may cause insulated portions of connector pin to swell and become difficult to insert into Pulse Generator.

- use soft rubber vessel loops to gently lift the nerve, if necessary.

Dr. Collins instead of vessel loop places small cottonoid underneath nerve.

In my experience, best is to place 1.0x0.5 inch piece of gelatin sponge (Gelfoam) soaked in thrombin solution – very soft, moist, and hemostatic; if small piece of sponge gets caught by instrument or gets trapped between the electrode and nerve, it is easy to pull it out and it is fast absorbable.

- two helical electrodes and one helical anchor tether are coiled around nerve.

- remember – negative electrode (with green suture embedded in helical material) goes cranial and positive electrode (with white suture) goes caudal:

N.B. if contacts are applied upside down, VNS will stimulate efferent fibers – may cause diarrhea!

- there is no preferred order how to place helices – either start with negative electrode first or start with anchor first (author’s preferred way).

- with forceps, gently pull each end of helical, using attached sutures to spread helical.

- starting with opened helical spread directly above and parallel to exposed nerve, turn helical clockwise at a 45 degree angle to nerve:

- place turn of helical where lead wire connects to helical (section with metal ribbon) onto nerve:

- using fine-tipped forceps, holding the suture ends, manipulate each end of helix to wrap around the nerve; usually it requires a coordinated effort of three or four hands:

Dr. Collins uses blunt nerve hook to wrap (procedure can be done by one person!) – stretches helical so metal contact sits on nerve; releases stretch so helical recoils and grasps nerve; then nerve hook goes under nerve, hooks and pools one side of pigtail; repeat on the other side of pigtail:

Source of picture: Viktoras Palys, MD >>

- form strain relief bend (lay it inside carotid sheath) and strain relief loop to provide adequate slack and allow for neck movement.

Source of pictures: Livanova >>

Dr. Collins never places tie-down tabs under the skin (risk of erosion!); he places tabs at least under platysma.

Pulse Generators

- do not use electrosurgical equipment after Pulse Generator has been introduced to sterile field - exposure to this equipment may damage Pulse Generator.

- Pulse Generator is implanted just below clavicle in subcutaneous pocket in the left upper chest along axillary border.

- place extra coiled lead to side of Pulse Generator, not behind it.

Do not place lead slack under Pulse Generator, because doing so could result in insulation failure and system malfunction.

- back out setscrew(s) and insert lead:

- it is ok to implant battery upside down.

- note for dual pin leads (now obsolete) - reversal of lead polarity has been associated with increased chance of bradycardia in animal studies.

Testing

- if patient has experienced asystole, severe bradycardia (heart rate < 40 bpm), or clinically significant change in heart rate during System Diagnostics (lead test) at time of initial device implantation, patient should be placed on cardiac monitor during initiation of stimulation.

models 100-102R, the test will stimulate at 1 mA for 14 seconds.

models 103-106, the test takes approximately 14 seconds; 1 mA output current (at 20 Hz, 500 µsec) will be used when the device is set to 0 mA (which is the case for new IPGs); otherwise, it will use the programmed output current (e.g. 0.75 mA for 5 seconds if programmed at 0.75 mA with ON time for 5 seconds, or 2.5 mA for 30 seconds if programmed at 2.5 mA with ON time for 30 seconds).

model 1000, the test uses the programmed output current for approximately 4 seconds.

N.B. have screw ready!!! – if patient goes into severe bradycardia, may quickly disconnect lead (as stimulator goes into testing cycle and won’t be easily stopped)

- two types of tests (interrogation):

System Diagnostics – if it fails (lead impedance “HIGH” or “LOW”) → Generator Diagnostics (test performed with test resistor that is included in Pulse Generator packaging or use accessory kit #502) - this test will verify that Pulse Generator is functioning properly, independent of lead.

N.B. IPG does not need to be in a pocket – impedances are checked between lead contacts (vs. in DBS, impedances are also checked between a contact and IPG case – so called “monopolar configuration”).

Source of picture: Livanova >>

- when implanting ≥ 106 model - make sure it senses heart rate (IPG has to be in a pocket to test this); select the lowest sensitivity level that correctly detects heart rate.

- sensitivity can be selected by averaging two R-wave amplitude measurements:

N.B. other stimulators (e.g. DBS) may interfere with proper heart rate detection.

Intraoperative programming

- see below >>

Closure

- all wounds are copiously irrigated with diluted povidone-iodine (Betadine) solution and filled with 1 gram of vancomycin powder.

- secure IPG by placing braided nonabsorbable suture (e.g. 2-0 silk) through suture hole and attaching it to underlying pectoralis fascia (not to muscle).

N.B. battery has only one hole for anchoring!

- deep interrupted absorbable (e.g. 2-0 Vicryl) stitches are placed to approximate the bulk of soft tissue (adipose tissue in the chest, platysma muscle in the neck).

- skin edges are sutured with interrupted inverted (deep dermal) stitches using thinner absorbable sutures (e.g. 3-0 or 4-0 Vicryl).

- if there is an increased risk of skin dehiscence, place additional simple running absorbable monofilament suture (e.g. 4-0 Monocryl) on epidermis for extra reinforcement.

- skin adhesive (e.g. Dermabond) is applied; once this had completely (!) dried, a sterile dressing (e.g. Tegaderm) is applied.

Postoperatively

- old school was not to program Pulse Generator to stimulation treatment for at least 14 days after initial implantation; now we start stim at lowest settings (0.25 mA) in OR.

- rarely, neck brace can be used for first week to help ensure proper lead stabilization. In contrary, most experts encourage range-of-motion neck exercises to minimize bowstring scar formation along the lead.

Patient Instructions

- postoperatively, observation in a hospital for a few hours after surgery is recommended, mainly for watch of neck hematoma that could compromise the airway.

  - make sure the patient has no vocal cord dysfunction but also educate that intermittent voice hoarseness is expected if VNS stimulation is turned on before discharge to home.

- Bathing: after discharge, patients are allowed to take a shower anytime, but prolonged soaking the incision is not allowed for 6 weeks.

- Dressing: patients are instructed to remove the dressing at 3 days after surgery, noting that the skin glue will stay on and flake off on its own in a few weeks.

- in rare cases, a neck brace can be used for the first week to help ensure proper lead stabilization. However, in contrast, most experts encourage gentle range-of-motion neck exercises for 6 weeks to minimize bowstring scar formation along the lead.

- patient should be reminded that VNS will not result in seizure cessation, thus, they should continue to practice seizure precautions and continue their antiseizure medications.

Source of picture: Viktoras Palys, MD >>

skin glue will stay on and will flake off on its own in a few weeks.

Regimen: for 6 weeks perform gentle neck range of motion exercises.

If, at any time, you notice signs of infection (worsening pain, redness, swelling, drainage), please contact the surgeon’s office ASAP.

MRI

Full instructions >>

Check vnstherapy.com/healthcare-professionals/mri – really easy to use website to find recommendations for your device!

- compatible with 1.5-3.0 T MRI (DBS is also approved for 3.0 T; RNS approved for 1.5 T).

Group A - 103, 105, 106, 1000 implanted higher than 4th rib (armpit level)

- OK to use transmit body coil (15 min scan, 15 min break time) and 3T MRI.

- if the patient requires a MRI of the C7-T8 area using a head/extremity coil or C7-L3 area using a body coil, surgical removal of the VNS is required.

Group B – 102, 104, also Group A implanted below 4th rib

- surgical removal of the VNS is required if MRI using a transmit RF body coil is needed.

- not all head RF coils are transmit and receive type; many are receive only; use of any local receive coil with the body coil in RF transmit mode presents the same RF heating hazards as the body coil alone with no local coils.

Cut / Broken Lead

- safe length of lead segment remaining (i.e., ≤ 2 cm) can be assessed by taking an x-ray. The length of 2 cm can be approximated by visualizing the distance between the positive and negative electrode (~1 cm). By design, there is approximately 1 cm between the positive electrode and the anchor tether, which is also likely remaining. Surgeons are instructed to remove as much of the lead as possible if explanting a system.

- length of lead segment remaining (i.e. ≤ 2 cm): MRI using the body coil, or MRI of the head or extremities with a head coil or local (extremity) coil (respectively) is allowed if the lead is transected here:

- if the lead is transected > 2 cm, only a transmit/receive head MRI or transmit/ receive extremity MRI is recommended; full body MRI is not allowed.

Warning: If it appears that > 2 cm of lead remains, then the patient cannot have an MRI with the body coil, but can still have an MRI using a extremity transmit/receive or head transmit/receive coil. Abandoned lead wires present increased risk of thermal injury to patients during MRI procedures based on their length and their exposure to RF.

Device preparation for MRI scan

- Perform an interrogation and record the information in the patient record (this information is used to restore the device settings after the MRI or in the rare case of a reset):

- Program the parameter settings as follows (VNS needs to be turned off by programmer, not by magnet):

Normal Output Current: 0 mA

Magnet Current: 0 mA

Model 106 and 1000 only:

Detection “OFF”

AutoStim Output Current: 0 mA

Turn off any other optional device features (Model 1000 only).

N.B. practical advice – only need to turn magnet to 0 mA (all other parameters can be left unchanged)

- Verify that placement of VNS is located between C7-T8.

Caution: MRI with a VNS implanted outside C7-T8 has not been evaluated!

Electrosurgery

If patient with implanted VNS is undergoing another surgery, it OK to use monopolar coagulator (Bovie) if far from VNS. Place grounding pad to avoid VNS system being in between grounding pad and monopolar coagulator use site

Programming

Stimulation waveform

Strength–duration time constant (chronaxie)

- as the duration of stimulus increases, the strength of the current required to activate a single fiber decreases.

Strength-duration relationship of the vagus nerve modeled by the Lapicque equation using experimentally collected nerve fiber recordings in canine

L. Lapicque. Definition experimentale de l'excitabilite. Soc Biol, 77 (1909), pp. 280-283.

C.D. Smith, L.A. Geddes, J.D. Bourland, K.S. Foster, W.E. Schoenlein. The chronaxie and propagation velocity of canine cervical vagus nerve fibers in vivo. Cardiovasc Eng: Int J, 1 (2) (2001), pp. 77-84

- model represents upper and lower bounds around 1.625 mA.

- minimum dose of VNS needed to activate afferent vagal fibers in the absence of glial scarring is derived from computational modeling

S.L. Helmers, et al. Application of a computational model of vagus nerve stimulation. Acta Neurol Scand, 126 (5) (Nov 2012), pp. 336-343

Canonical biophysical laws

Dosing a VNS system can be simplified functionally to a two-step process:

  - activate the nerve with current and pulse width (the ‘volume’) - output current and pulse width work in concert to activate neural tissues.

  - modulate central nuclei via temporal code (the ‘message’) - therapeutic ‘message’ of neural modulation depends on the signal frequency and duty cycle.

Intraoperative

(e.g. in OR or PACU right after VNS implantation):

- Interrogate generator

- Perform System Diagnostics (system stimulates at 1 mA or set output current – watch for heart rate change)

- Program – normal mode:

output current 0.25 mA >>

frequency 20 Hz >>

pulse width 250 msec >>

on time 30 seconds, off time 5 minutes

- Program – magnet mode: current 0.25 mA higher than output, 250 msec pulse width for 30 seconds

- Program – AutoStim mode:

Verify Heartbeat Detection (check is device detected HR and pulse oximetry HR match)

heart rate threshold 50%

current 0.125 mA higher than normal output

stim time 30 secs

- Always Interrogate generator again as last step in session to verify settings.

N.B. only increase current from factory 0 mA to 0.25-0.5 mA; for the rest – may leave factory defaults.

Magnet Mode

- patient (or caregiver) can use magnet (e.g. in bracelet) to give additional boost stimulation PRN (program 0.25 mA higher than regular mode with 250-500 msec pulse width for 30* seconds).

*default is 60 secs but in studies seizures stop within 10 seconds, thus, 30 seconds is plenty.

- if seizures are successfully aborted with magnet, increase regular mode current to the one that magnet uses.

- if magnet is held over device for > 3 seconds, device is turned off and stays off as long as magnet is held on (sometimes patients make shirt pocket to hold magnet to keep magnet on and VNS off – marathon runners so VNS does not interfere with breathing; singers – so voice does not change).

AutoStim – Tachycardia Detection

82% of patients have ictal tachycardia (if epilepsy involves autonomic centers; cf. frontal lobe epilepsy usually has no tachycardia) – candidates for AutoStim.

AutoStim - detects and responds to heart rate increases that may be associated with seizures (closed-loop system)

N.B. just sensing heart rate consumes approx. 20% (2 years) of battery life! – choose carefully!

- AutoStim monitors the last 5 mins of heart rate to have baseline rate.

- AutoStim triggers if heart rate increases from baseline to threshold over 10 seconds (vs. during physical exercises, heart rate increases slower).

- why not to use just AutoStim and avoid chronic stimulation (thus saving battery):

- longterm benefits (seizure control, depression treatment) depend on chronic stimulation.

- vagus nerve is not used to receive stimulation, so AutoStim causes “shock” sensation.

- heart rate detection may not work due to presence of electrical artifacts from other stimulators (e.g. DBS).

Ictal tachycardia

- true ictal tachycardia (> 100 bpm with at least 55% increase or 35 bpm increase from baseline) is observed in 16%-17% of seizures; when using a more liberal threshold of ≥ 20% increase in heart rate, up to 66% of seizures could be detected but at the cost of 7 false detections per hour.

Programming

- Threshold for AutoStim: set initial heart rate threshold 50% (setting range is from 20% to 70%. 20% is most sensitive. 70% is least sensitive).

- Heartbeat Detection (Sensitivity) ranging from 1 to 5, with “1” being the least sensitive and “5” being the most sensitive setting.

N.B. model 106 is only capable of detecting heart beats in the range 32-240 bpm (vs. model 1000 - 28-180 bpm)

- program 0.125 mA higher than normal mode (i.e. in the middle between chronic stim and magnet mode) for 30* seconds.

*in studies, seizures stop within 10 seconds, thus, 30 seconds is plenty, plus, resets heart rate calculator sooner (vs. if stim was for 60 secs).

- to have AutoStim enabled, off periods have to be > 0.8 min.

Studies

The multicenter European Cardiac Based Seizure Detection Trial (E-36) - tested a heart rate–triggered VNS device in 30 patients with frequent seizures and well-established drug-resistant epilepsy that was not amenable to epilepsy surgery. They were also chosen on the basis of frequent tachycardia, the feature that the device was designed to detect.

Bradycardia Detection – “SUDEP monitor”

- enabled only for 7.5 mins after magnet swipe or AutoStim (i.e. VNS monitors if patient goes into bradycardia during or after seizure – SUDEP risk).

Impedance

- normal impedances are < 2000-4000.

  - high lead impedance is defined as ≥ 5300 Ohms, however, there is no upper limit where IPG will stop delivering therapy (all depends on settings; if impedance > 10 kOhm, likely there is very little therapy being delivered).

  - low lead impedance is defined as ≤ 600 Ohms.

Relationship of delivered output current (mA) to lead impedance for 1000 µsec pulse with output currents from 0 to 3.5 mA:

- if battery is dead

  - preop: cannot test impedances (but starting with model 103, battery retains charge for lead testing).

  - intraop: VNS does not have external extension leads to test impedance if battery is dead; thus, need to connect new battery intraop.

Caveat: if using new 104 or 102R and found that dual pin lead is broken, will need to replace to the new single pin lead [LivaNova does not manufacture new dual-pin leads for replacement], you will need new single-pin battery [LivaNova does not manufacture dual-pin-to-single-pin adaptors] = you wasted new 104 or 102R battery!!!!

Clinic - Programming Phase I

aka “1 tab BID → 2 tab BID”

- gradually increase of output current - stepwise, every 2 weeks*, up to 2.5 mA** (as long as patient is seeing improvement; when plateaus, go one step back and start phase II).

*can be done every 15 mins in clinic if patient tolerates side effects

**1.5 mA should activate all axons so no real need to go beyond that (all further efficacy depends on duty cycle) but there are patients that benefit from 3.5 mA.

- if VNS is implanted for emergency cases (patient in status), current may be ramped up quickly.

Output Current

- based on Fahoum et al. 2022 study – target current is 1.625 mA at 250 μsec

  - increasing output current > 1.625 mA will not further improve seizure control in patients that have already demonstrated response to VNS.

  - increasing output current > 1.625 mA is unlikely to convert a non-responder into a responder,

  - children tend to need higher mA than adults.

Pulse Width

- 250 msec - all axons get excited (default is 500 msec but that increases side effects, e.g. sleep apnea, and drains battery).

- there is little functional reason to utilize pulse widths > 250 μsec for clinical purposes due to battery considerations.

- if it is clinically necessary to use other pulse widths than 250 μsec, one should consider increasing or decreasing output current consistent with the vagal strength-duration relationship.

Frequency

- assuming all VNS fibers are activated (therapeutic combination of output current and pulse width), the final dose parameter is Signal Frequency.

- ascending nuclei in CNS are responsible for decoding this message and may change their activity according to temporal and rate components.

- animal studies indicated low frequencies (20 Hz–30 Hz) have a stronger anticonvulsive effect.

- migration toward 20 Hz from the original manufacturer setting of 30 Hz occurred in response to anecdotal reports of improved mood effects at 20 Hz, without impact to anti-convulsive effect (LivaNova changed factory setting from 30 Hz to 20 Hz with the introduction of the most recent pulse generator).

Best VNS dose

Fahoum, F. et al. VNS parameters for clinical response in Epilepsy. Brain stimulation, Volume 15, Issue 3, May–June 2022, Pages 814-821

- retrospective analysis of 1178 subjects from twelve LivaNova-sponsored clinical studies.

N.B. studies included interventional and observational designs, different follow-up durations, targeted patients of slightly different demographic profiles, and the methods for data collection were not uniform across all studies; none of these studies were prospectively designed!

- patients titrated near 1.5-1.75 mA experienced greater seizure frequency reduction than those titrated above or below it.

VNS response may be maximized near 1.625 mA [1.61 mA*], 30 sec ON, 3 min OFF [17.1%*]

*based on study model

- study did not find that any of the pulse widths or signal frequencies had a significant impact on response.

VNS titrated to settings near the model-selected dose of 1.625mA (any pulse width) → response rate was calculated at 12 months ater implant:

  - presence of a peak probability of response implies that patients programmed to higher VNS intensities may not experience further improvements in efficacy, and some patients may not respond to VNS at any intensity.

Generalized Linear Mixed Model (GLMM) to identify relationships between both VNS parameters and demographic features to clinical response (≥ 50% reduction in seizures). This is a population-level outcome, indicating that some patients will require VNS intensities above or below this value.

| Population level mean and standard deviation of outcomes predicted by the GLMM (probability of response at 12 months for a patient who is 26 years old with epilepsy duration of 15 years and duty cycle set at 10.3%): | GLMM predicted outcome for several patients for a random sample of hypothetical 50 patients with specific demographic parameters (for each curve the probability is maximized at 1.61 mA of output current): |
| --- | --- |

Outcomes of patients after 12 months who were initially titrated* to the target intensity of VNS (1.5 mA–1.75 mA) and who were responders at 12 months (n = 186 subjects)

*but may have increased or decreased their VNS dose at follow up visits after 12 months

  - analysis suggests that patients with initial response to VNS are highly likely to retain their clinical response over time and maintain clinically significant reductions in seizure frequency.

  - analysis demonstrates a plateau effect, whereby patients titrated to output currents higher than the proposed target output current did not achieve greater likelihood of response or greater reductions in seizure frequency.

  - identical analysis for target duty cycle obtained similar results where further increases in duty cycle did not confer additional higher response rates or reductions in seizure frequency.

There was no indication in this long-term follow up that increasing output current resulted in higher likelihood of response or further improvements in seizure frequency.

Therefore, we conclude that use of output currents well beyond the target population level settings of 1.625 mA at 250 μsec is unlikely to convert a non-responder into a responder, and unlikely to offer superior levels of response in patients that have already demonstrated response to VNS.

Clinic - Programming Phase II

aka “2 tab BID → 2 tab TID”

  - duty cycle should be increased only if patients do not respond at 10% duty in favor of battery savings.

  - patients who do not respond at the introductory duty cycle of 10% should have their duty cycle increased.

- increasing duty cycle every 3 months.

- side effects do not get worse during phase II.

- based on Fahoum et al. 2022 study – target duty cycle is 17.8% is associated with the highest probability of response.

- “rapid cycling” - higher duty cycle (definition varies from OFF-Time ≤ 1.8 min to a combination of ON-Time = 7 s and OFF-Time = 0.2–0.3 min) in case standard cycling is found to be not effective in a patient; safe and potentially more effective than the standard cycling in pediatric patients.

Duty Cycle

- stimulation (on time) as a percent of the whole time.

- no studies on how many hrs/day are needed to achieve neuromodulatory effect (based on animal studies, 100% [i.e. continuous stim] would be the best but depletes battery fast).

duty cycles in gray are not recommended as they represent parameter combinations with ON > OFF time

N.B. high frequencies and duty cycles used in combination have been previously reported to damage neural tissue in animals → LivaNova cautions against programming duty cycles > 50%, even at lower frequencies, in favor of safety.

D.R. Merrill, M. Bikson, J.G. Jefferys. Electrical stimulation of excitable tissue: design of efficacious and safe protocols. J Neurosci Methods, 141 (2) (Feb 15 2005), pp. 171-198

Side Effects

VNS does not carry the CNS side effects associated with oral ASMs!

  - VNS usually leads to improvement of behavioral problems and has mood stabilizing effects.

  - acute psychosis and depression have been described in relationship to VNS.

  - effects on memory are variable: some authors report improvement in word-recognition memory, others have found decline in figural memory during VNS.

  - central pain suppression has been reported with VNS.

- hoarseness or change in voice (30-50%)

- cough, paresthesia, pain, dyspnea, or headache

- sleep apnea (in adults, esp. if 500 msec pulse width is used).

Significant risk of de-novo or aggravating sleep breathing disorders in up to 57% of patients!

- asystole (reports of occurring > 1 year after VNS surgery)

- swallow studies in 8 children with VNS did not show any evidence of aspiration (only one patient had laryngeal penetration of barium caused by VNS, but no aspiration was seen).

- increased drooling as a side effect of VNS has been reported in patients with Lennox-Gastaut syndrome.

Generator data communication produces an ECG artifact:

Complications

- surgical bed hematoma (in the neck can cause airway compromise) – main reason patients are kept for observation for several hours postop – in extremis, may require bedside neck exploration.

- infection (3-8%); when identified early, antibiotic therapy should be initiated with the hopes of preventing device removal. Often however, the entire system needs to be explanted until the infection is cleared. Once cleared, VNS syetm can be re-implanted with consideration to use antibiotic-impregnated absorbable pouch for IPG pocket.

Smyth M, Tubbs S, Bebin E, et al. Complications of chronic vagus nerve stimulation for epilepsy in children. J. Neurosurg 2003;99:500–3.

Patel N, Edwards M. Vagal nerve stimulator pocket infections. Ped Inf Dis J 2004;23:681–3.

Ortler M, Luef G, Kofler A et al. Deep wound infections after vagus nerve stimulator implantation: treatment without removal of the device. Epilepsia 2001;42:133–5.

Murphy J, Hornig G, Schallert G et al. Adverse events in children receiving intermittent left vagal nerve stimulation. Pediatr Neurol 1998;19:42–4.

- there are no long term side effects of chronic stimulation (reported change in axon diameter but no physiologic effects); however, excessive stimulation has resulted in degenerative nerve damage in laboratory animals, esp. when ON time exceeds OFF time (duty cycle > 50%), which can be produced by continuous or frequent magnet activation (> 8 hours).

  - EMG of the cricothyroid and thyroarytenoid muscles indicates that degeneration of the recurrent laryngeal nerve occurs in 60% of patients after prolonged repetitive stimulation.

Shaffer MJ, Jackson CE, Szabo CA, Simpson CB. Vagal nerve stimulation: clinical and electrophysiological effects on vocal fold function. Ann Otol Rhinol Laryngol 2005;114(1Pt1):7–14

- damage to surrounding structures leading to vocal cord paralysis, Horner’s syndrome, lower facial muscle paresis (likely due to incisions too high in the neck).

- cardiac arrest.

- hardware failure - lead fracture, generator malfunction (2.7% of 74 patients with a minimum follow up of 1 year).

Smyth M, Tubbs S, Bebin E, et al. Complications of chronic vagus nerve stimulation for epilepsy in children. J. Neurosurg 2003;99:500–3.

Drug interactions

- best results when combined with SV2A modulators (levetiracetam, brivaracetam), slow Na channel inhibitors (eslicarbazepine).

- VNS combined with lamotrigine can induce mania (esp. pt with hx of psych, esp. bipolar).

- GABAergics (valproate and gabapentin) can worsen sleep apnea - do not combine with VNS, esp. in obese.

Outcomes - Depression

- 1-year response rate 30-53% (vs. 10% with best non-VNS therapies).

- from an FDA-mandated registry (Aaronson, 795 patients with TRD and 5-year follow-up): VNS plus treatment as usual vs. treatment as usual had higher cumulative response rates (68% vs 41%) and remission rates (43% vs 26%).

Observational registry study

Scott T. Aaronson et al. A 5-Year Observational Study of Patients With Treatment-Resistant Depression Treated With Vagus Nerve Stimulation or Treatment as Usual: Comparison of Response, Remission, and Suicidality. Am J Psych. Published Online:31 Mar 2017 >>

- 5-year, prospective, open-label, nonrandomized, observational registry study

- 61 U.S. sites, 795 patients.

- inclusion - treatment-resistant depression:

- major depressive episode (unipolar or bipolar depression) of at least 2 years’ duration

- ≥ 3 depressive episodes (including the current episode)

+ failed ≥ 4 depression treatments (including ECT).

- exclusion - history of psychosis or rapid-cycling bipolar disorder.

- primary efficacy measure - response rate = decrease of ≥ 50% in baseline Montgomery-Åsberg Depression Rating Scale (MADRS) score at any postbaseline visit during the 5-year study.

- secondary efficacy measure - remission.

- results - VNS group had better clinical outcomes:

  - 5-year cumulative response rate 67.6% (vs. 40.9%).

  - cumulative first-time remitters 43.3% (vs. 25.7%).

  - subanalysis of history of response to ECT:

    - positive response to ECT- 5-year cumulative response rate 71.3% (vs. 56.9%).

    - no response to ECT - 59.6% (vs. 34.1%).

- conclusions - adjunctive VNS has enhanced antidepressant effects compared with treatment as usual.

Pending read

Outcomes - Epilepsy

There are > 150,000 patients with VNS but still we don't know who are best candidates - because patients were not properly worked up preop.

1/3 VNS pts are nonresponders

- because of its lower risk for neurologic injury – VNS was a standard by which all other more invasive therapies must be judged.

- long-term stimulation (> 6 months) → greater rather than lesser effect, i.e. VNS has improved efficacy over time!!!

- 50% reduction in seizure frequency in 50-60% of patients after 2 years of therapy (76% at 8 years)

- up to 8% patients become seizure free.

- 25% of patients with long-term f/u have no benefit!?

- it remains unclear if VNS allows to decrease ASMs.

- at 12 months, VNS improves quality of life by 4.5-fold (PULSE study 2014).

- based on Fahoum et al. 2022 study:

  - older people are more likely to achieve response (vs. pediatric patients);

  - less delay between epilepsy diagnosis and eventual VNS - more likely to achieve response.

Example of consult:

“VNS therapy has a 60% chance of a 50% reduction in seizures and low risks of complications, but they do include infection (3-4%), hoarseness (up to 50%), neck hematoma, cardiac arrhythmia, even coma and death. It would not cure the epilepsy, thus, VNS is not indicated if the patient’s sole goal is to eliminate seizures and start driving. During appointment approximately 2 weeks after the surgery, the device will be activated but further programming will be done by epileptologist.”

E03 trial (multinational)

E05 trial (US only)

CORE-VNS (NCT03529045)

- Comprehensive Outcomes Registry in Subjects With Epilepsy Treated With Vagus Nerve Stimulation Therapy

Outcomes by VNS Registry + Meta-analysis

VNS therapy Patient Outcome Registry (5554 patients) + systematic review of the literature (2869 patients across 78 studies)

Englot et al. “Rates and Predictors of Seizure Freedom With Vagus Nerve Stimulation for Intractable Epilepsy” Neurosurgery: September 2016 - Volume 79 - Issue 3 - p 345–353

| Duration | responder rate – VNS registry% (review%) | seizure freedom (Engel Ia) – VNS registry% (review%) |
| --- | --- | --- |
| 0-4 mo | 49% (40%) | 5.1% (2.6%) |
| 24-48 mo | 63% (60.1%) | 8.2% (8%) |

- on multivariate analysis, seizure freedom was predicted by age of epilepsy onset >12 years (odds ratio [OR], 1.89; 95% confidence interval [CI], 1.38-2.58), and predominantly generalized seizure type (OR, 1.36; 95% CI, 1.01-1.82), while overall response to VNS was predicted by nonlesional epilepsy (OR, 1.38; 95% CI, 1.06-1.81).

VNS therapy registry:

- All patients - progressive increase in seizure freedom, paralleling increases in the rate of response to therapy (≥ 50% seizure frequency reduction) and the median reduction of seizure frequency:

- Partial seizures:

- Generalized seizures:

- seizure freedom was significantly more likely in generalized seizures at 0 to 4 months (P < .01, Pearson χ2) and at 4 to 12 months of follow-up, although this difference was not significant at 12 to 24 months or 24 to 48 months (P > .5).

- responder rate and median seizure reduction did not differ significantly between patients with primarily partial vs generalized seizures.

Seizure freedom rates by age of epilepsy onset – older-age onset patients do better:

Seizure freedom rates by age of implantation - no significant difference in age of implantation was observed between patients with or without seizure freedom (P > .2 at each time point, Wilcoxon sum rank test):

Seizure freedom rates by onset-to-implant interval - patients who achieved seizure freedom showed shorter onset-to-implant intervals than those with persistent seizures, but this difference was not significant (P range 0.07-0.26 at each time point, Wilcoxon sum rank test):

Systematic literature review:

Progressive increase in both seizure freedom and response rate over time, resembling findings from the registry.

- at latest follow-up, 60% of patients responded to VNS (≥ 50% reduction in seizures), and 8% were seizure free:

Pediatric VNS outcomes

LoPresti, Melissa et al. Pediatric Vagus Nerve Stimulation: Case Series Outcomes and Future Directions. Neurosurgery 92(5):p 1043-1051, May 2023.

- single center (Texas Children's Hospital), retrospective study, 297 patients.

- 78% of patients showed at least a partial reduction in the frequency of their total number of seizures, with 52% achieving at least 50% reduction and 11% of patients cured of main seizure subtype at the last follow-up.

- there was no correlation between main seizure reduction and age at diagnosis, age at surgery, or duration of symptoms, genetic vs non-genetic, congenital vs non-congenital.

- predictors of VNS outcome:

- MRI findings:

- brain atrophy → significantly associated with worse VNS outcomes.

- dysplastic hippocampus and chronic periventricular leukomalacia → improved VNS outcomes.

- patients with more variable disease (greater number of seizure types, more diverse semiology) → improved VNS outcomes.

- patients with a higher preoperative seizure burden tended to benefit more (r = 0.15, P = .045 Spearman ρ).

- small positive trend between duration of follow-up and total seizure reduction (r = 0.13), suggesting that outcomes may improve over time, but this did not reach significance (P = .1)

Different epilepsy types

Meta-analysis suggests VNS is more effective in generalized seizures than partial seizures:

Englot et al., J Neurosurg Dec 2011

Generalized Epilepsy

| Study | Response | Follow up |
| --- | --- | --- |
| Labar (E04) | 11/24 (45.8%) | 3 mo |
| Holmes (U Wash) | 7/16 (43.8%) | 12-21 mo |
| Ng (Mt Sinai) | 14/27 (51.9%) | Av 21.6 mo |
| Kostov (Oslo) | 20/30 (66.7%) | Av 52 mo (17-123) |

Symptomatic epilepsy

Tonic clonic  N=14              -33%

Atypical absence   N=8        -48%

Tonic  N=4                           -70%

Myoclonic   N=3                  -80%

Idiopathic epilepsy

Tonic clonic   N=4               -76%

Typical absence  N=4           -81%

Myoclonic  N=3                     0%

Labar, D et al. Vagus nerve stimulation for medication-resistant generalized epilepsy. Neurology. 52(7):1510-1512, April 22, 1999

Hippocampal Epilepsy

- Alsaadi et al (UCSF) – 6/10 pts (60%) with > 50% seizure reduction at 1 yr.

- Kuba et al (Brno, Czech) – 3/8 (37.5%) at 1 yr and 5/8 (62.5%) at 18 mo.

Tumor-associated epilepsy

VNS in tumor associated medically intractable epilepsy

Efficacy of vagus nerve stimulation as a treatment for medically intractable epilepsy in brain tumor patients. A case-controlled study using the VNS therapy Patient Outcome Registry. Seizure - European Journal of Epilepsy — Patel KS, et al. June 17, 2013

VNS is a viable option assuming cytoreductive and other adjuvant therapies have been fully explored.

- data from the VNS therapy Patient Outcome Registry.

- 107 patients with an epilepsy etiology related to a brain tumor, responder rate of 48% at 3 months and 79% at 24 months – similar to non-oncological patients (no statistical difference in seizure reduction compared with 326 case–control patients from the registry without brain tumors)

- no significant difference in AED usage from baseline to 24 months post implant in either group.

Cost-utility analysis

Pediatric patients

Helmers SL, Duh MS, Guerin A, et al. Clinical outcomes, quality of life, and costs associated with implantation of vagus nerve stimulation therapy in pediatric patients with drug-resistant epilepsy. Eur J Paediat Neurol. 2012;16(5):449–458.

- retrospective analysis using Medicaid data

- patients 1-11 years old (N = 238):

- hospitalizations and ER visits were reduced Post-VNS vs. Pre-VNS (adjusted IRR = 0.73 [95% CI: 0.61-0.88] and 0.74 [95% CI: 0.65-0.83], respectively).

- average total healthcare costs were lower Post-VNS vs. Pre-VNS ($18,437 vs. $18,839 quarterly [adjusted p = 0.052]).

- lifetime QALY gain after VNS was 5.96 years.

- patients 12-17 years old (N = 207):

- hospitalizations and status epilepticus events were reduced Post-VNS vs. Pre-VNS (adjusted IRR = 0.43 [95% CI: 0.34-0.54] and 0.25 [95% CI: 0.16-0.39], respectively).

- average total healthcare costs were lower Post-VNS vs. Pre-VNS period ($14,546 vs. $19,695 quarterly [adjusted p = 0.002]).

- lifetime QALY gain after VNS was 4.82 years.

Model 100 VNS

Ben-Menachem E, Hellstrom K, Verstappen D. Analysis of direct hospital costs before and 18 months after treatment with vagus nerve stimulation therapy in 43 patients. Neurology. 2002;59(6 suppl 4):S44–S47.

- therapy response was defined as ≥ 25% reduction in seizure frequency.

- aim was to assess unplanned hospital costs 18 months before and 18 months after VNS implantation in 43 patients (Sahlgrenska University Hospital in Sweden) - irrespective of whether they responded to VNS therapy!

- for all patients, ICU costs were reduced from 46,875 to 0 US dollars, ER visits from 13,000 to 9,000 US dollars, and ward stays from 151,125 to 21,375 US dollars.

- total hospital costs (for all 43 patients) before VNS therapy were 211,000 US dollars and after 18 months of treatment were reduced to 30,375 US dollars = average cost savings of ≈ 3,000 USD/year/patient in unplanned hospital costs.

Bibliography for ch. “Epilepsy and Seizures” → follow this link

Viktor’s Notes℠ for the Neurosurgery Resident

Please visit website at www.NeurosurgeryResident.net