# Responsive Neurostimulation (RNS)

Last updated: September 18, 2024

Brochures	2

Philosophy	2

Indications	2

FDA approved (2013)	2

Status Epilepticus	3

Contraindications	3

Pathophysiology	4

Hardware	4

Device	5

300, 320	5

Poseidon	5

Strips	6

Depths	6

Burr hole cover	7

On imaging	8

Preop counselling	9

Strategies	9

EcoG via RNS	9

Neocortical Broad Onsets	9

Concurrent Partial resection + RNS	12

Bilateral Mesial Temporal Lobe Epilepsy	13

Thalamic implantations	15

Procedure – Depth Electrode implantation	15

Hippocampal depth	15

Cannula options	15

Incisions	16

Procedure - Device Implantation	20

Craniectomy	20

Leads	20

Procedure - Battery Replacement	21

Battery charge level	21

Procedure - Explantation	21

Complications	21

Postoperatively	21

MRI	21

Contraindicated procedures	22

Analysis & Programming	23

Detections	24

Stimulation	25

Outcomes	26

Neuropsychological	26

Quality of Life	26

Seizures	27

Studies	27

SUDEP	31

Medical Treatment Monitoring	32

Seizure prediction	33

RNS for Thalamus – see p. E27 >>

Comparison of Neuromodulations (RNS, DBS, VNS) – see p. E11 >>

Brochures

- see >>

Reading
E. Geller. Responsive neurostimulation: Review of clinical trials and insights into focal epilepsy. Epilepsy Behav, 88 (2018), pp. 11-20

O. Devinsky, D. Friedman, R.B. Duckrow, N.B. Fountain, R.P. Gwinn, J.W. Leiphart. Sudden unexpected death in epilepsy in patients treated with brain-responsive neurostimulation. Epilepsia, 59 (2018), pp. 555-561

Philosophy

RNS vs DBS – see p. E11 >>

RNS = responsive neurostimulation on demand (closed loop) + diagnostic modality with recording (ambulatory ECoG).

Advent of RNS decreased tolerance to functional consequences of epilepsy surgery.

Indications

FDA approved (2013)

- adjunctive therapy in reducing frequency of seizures in individuals ≥ 18 years**** with partial onset seizures who have undergone diagnostic testing that localized no more than 2* epileptogenic foci, are refractory to ≥ 2 antiepileptic medications, and currently have frequent** and disabling*** seizures.

*Neuropace has receptacle only for two leads.

** ≥ 3 disabling seizures per month over three most recent months (with no month with fewer than two seizures); RNS® System has not been evaluated in patients with less frequent seizures.

***motor partial seizures, complex partial seizures and / or secondarily generalized seizures

****FDA indication for RNS is age > 18 years, but it likely has efficacy in adolescents and children (limitation in young children is skull thickness before it reaches adult size, so it could certainly be used in adolescents).

Summary

Focal modulations:

- ≤ 2 foci

- eloquent areas

- mesial temporal lobe (dominant uni- or bilateral)

- difficult to resect (e.g. insula, large regional onsets, interhemispheric)

- failed previous surgery or VNS (24% RNS patients have VNS)

Most modern indication – implant in thalamus for network modulation in generalized epilepsies (such as idiopathic [NAUTILUS study – centromedian nucleus of thalamus], Lennox-Gastaut [LGS Feasibility Study – centromedian nucleus of thalamus vs. ESTEL trial – DBS for centromedian nucleus of thalamus), multifocal epilepsies, epilepsies with regional onsets. see p. E27 >>

VNS, RNS, and DBS are all palliative and comparable in efficacy, both in pivotal trials and over longer-term trials. VNS is sometimes a first choice as it is extracranial. A specific scenario where RNS may have an advantage is bilateral mesiotemporal epilepsy – RNS allows for long-term ECoG recording, which may in turn (occasionally) allow for an eventual resection (of seizure-dominant hippocampus or prove that “bilateral” disease is de facto unilateral) in a small number of patients.

Network strategies of RNS

Richardson 2023

Status Epilepticus

Case reports of efficacy in refractory status epilepticus:

Source of picture: Dr. Kathryn Holloway (VCU)

Contraindications

From RNS® System manual:

- high risk for surgical complications such as active systemic infection, coagulation disorders (such as the use of anti-thrombotic therapies) or platelet count < 50,000.

- patients who have medical devices implanted that deliver electrical energy to the brain.

- patients who are unable, or do not have the necessary assistance, to properly operate the NeuroPace® Remote Monitor or magnet.

Pathophysiology

See p. E11 >>

Hardware

- RNS Neurostimulator Kit Model # RNS-300M-K.

- four cortical 5 mm bone screws.

Device

300, 320

- accepts two electrode arrays; if uncertain – may implant more (“sleeping soldier” lead) and leave contact end capped (may switch and reconnect later).

| Feature | RNS-320-K | RNS-300M-K |
| --- | --- | --- |
| Battery life | 8.4 yrs (under “medium” settings) | 3.9 yrs |
| Data capacity | 1 MB - stores 8 ECoGs (90 sec duration each)* | 0.5 MB |
| Price | 6000 USD more expensive (31,950 USD 35,250 USD)** | 25,950 USD |

*still no warning to the patient when memory is full and data is overwritten; why just 1 MB? – per RNS engineers, it is the best bytes / memory consumption ratio

**includes 1x RNS Neurostimulator, 1x Connector Cover, 1x Ferrule, 1x Ferrule Clamp, 1x Lead Strain Relief, 1x Torque Driver

Q: who should be getting model 300? A: nobody.

Kit includes also ferrule (675  745 USD):

Kit does not include craniectomy template (375  420 USD):

Practical tip – use nonsterile template (from rep) to plan craniotomy scalp incision;

for craniectomy – use ferrule to draw lines on skull.

Poseidon

- same battery capacity

- same size of memory (1 MB)

- connects to four leads

- backwards compatible

- system will have AI assistance

Strips

- 4100  4530 USD (includes 1x Cortical Strip Lead, 2x Lead Cap, 4x Suture Sleeve, 1x Tunneling Tool Straw)

- hand-made.

- 4 contacts spaced 10 millimeters apart - spans 33 mm

- lead lengths (tip - beware unnecessarily long leads):

15 cm (model CL-315-10)

25 cm (model CL-325-10)

35 cm (model CL-335-10)

Depths

- 4100  4530 USD (includes 1x Depth Lead, 2x Lead Cap, 4x Suture Sleeve, 1x Stop Gauge, 1x Tunneling Tool Straw)

- hand-made.

- 4 contacts

- diameter 1.27 mm (same as older DBS electrodes – so could use Stimloc burr hole device); slotted cannula 2.1 mm or DBS cannula (inner diameter 1.6 mm)

N.B. best to use DBS equipment for targeting to achieve best accuracy (esp. in z axis).

- two configurations (tip - beware unnecessarily long leads; too short leads maybe hard to secure to lead holder during implantation):

3.5 mm contact spacing, 30 cm length (model DL-330-3.5) – electrode span 12.5 mm

3.5 mm electrode spacing, 44 cm length (model DL-344-3.5) – electrode span 12.5 mm

10 mm electrode spacing, 30 cm length (model DL-330-10) – electrode span 32 mm

10 mm electrode spacing, 44 cm length (model DL-344-10) – electrode span 32 mm

Particularly useful for small targets (such as thalamic implantation)

Wider spacing is for hippocampal depths!

Burr hole cover

- 325  360 USD

- requires three 1.5 – 1.8 mm screws.

- there are other alternatives – see below >>

On imaging

Source of picture: Viktoras Palys, MD >>

Preop counselling

I explained the alternative of the VNS, pros and cons of each as well as the risks of surgery to include no change or worsening of seizures, visual fields compromise, memory compromise, or other injury to the brain.  I explained that how the computer was placed to replace the skull and the electrodes that pass through the brain to the hippocampus.  We talked about shaving some of the head. The patient understands the risk of an infection, which would require removal.

Strategies

EcoG via RNS

- may use RNS strip electrodes for live ECoG intraop but only if connected to IPG (it has to be grounded); RNS electrodes do not have adapters to connect to conventional ECoG machine.

Neocortical Broad Onsets

Neocortical seizure onsets localized to one “regional” (broad) area of cortex are difficult to treat with resective surgery.

- RNS may encompass 8-11 cm in regional stimulation.

- electrode contacts have to be on EZ else detection may not work well.

Strategy #1 – strip electrodes

Case series

Brandy Ma, Emily Mirro, Robert Knowlton, Edward Chang, Vikram Rao

- 12 patients who had one “regional” seizure onset zone were treated with the RNS System at UCSF.

- two depth and/or cortical strip leads were on average 3.3 cm apart (range: 1.2 cm - 6.0 cm).

- 80% median seizure reduction and 83% responder rate was achieved.

Strategy #2 – depth electrodes into neocortex

- can be preferred strategy for neocortical seizure onsets when cortical strip leads are not feasible or the localization was done with SEEG.

Case series

Allen Ho, Jonathon Parker, Emily Mirro, Babak Razavi, Kai Miller, Dora Hermes, Matthew Markert, Gerald Grant, Jaimie Henderson, Casey Halpern. Anatomic early electrographic seizure onset localization and clinical response with responsive neurostimulation of depth leads placed in neocortex.

- 16 patients who had ≥ 6 months follow-up had a median seizure frequency reduction of 65%.

Strategy #3 – depth electrode(s) into thalamus (e.g. centromedian nucleus)

- for generalized epilepsies (such as Lennox-Gastaut).

Concurrent Partial resection + RNS

RNS can be used concurrently with a resection such that treatment is extended to an unresectable area of the seizure focus.

Case series

Babak Razavi, Emily Mirro, Lawrence Shuer

- 3 patients at Stanford had a concurrent resection with RNS placement.

- all 3 patients had electrographic and clinical seizures after resection.

- 2 patients became seizure free and the 3rd patient reported 62.5% reduction and less severe seizures.

Bilateral Mesial Temporal Lobe Epilepsy

Long-term ambulatory ECoG monitoring of bilateral mesial temporal lobes with the RNS System can potentially lead to an efficacious resective or ablative procedure

Pre-RNS presumed lateralization changes in 20% of patients after chronic ambulatory ECoG:

| 11 patients presumed unilateral; 7/11 (64%) had bilateral electrographic seizures | 71 patients presumed bilateral; 9/71 (13%) had only unilateral electrographic seizures |
| --- | --- |

King-Stephens et al., Epilepsia 2015

Time to record first contralateral MTL seizure in patient with bilateral implants:

King-Stephens et al., Epilepsia 2015

Case series

Lawrence Hirsch et al. Outcomes after mesial temporal lobe resection following long-term ambulatory recording by the RNS System

- after bilateral MTL monitoring with RNS, 21 patients (at 17 epilepsy centers) had MTL resections.

- average RNS time before resection was 42 months (range: 8-117 months), and average post-resection follow-up was 37 months (range: 7-101 months).

- mean seizure reduction for all 21 patients was 90% (range: 50-100%).

- 15 of 21 patients (71%) were free of clinically reported disabling seizures.

- 8 of the 15 seizure free patients had only unilateral seizure onsets stored by the RNS System prior to resection, and the remaining 7 had bilateral seizures with an average of 90% from the resected side.

- 19/21 patients continued to be treated with the RNS.

Thalamic implantations

For thalamic targets – see p. E27 >>

Cameron P. Beaudreault et al. Responsive Neurostimulation Targeting the Anterior, Centromedian and Pulvinar Thalamic Nuclei and the Detection of Electrographic Seizures in Pediatric and Young Adult Patients. Front Hum Neurosci. 2022; 16: 876204. Published online 2022 Apr 12. doi: 10.3389/fnhum.2022.876204 PMCID: PMC9039390 PMID: 35496067

- bilateral thalamus

- unilateral thalamus + cortical strip (“corticothalamic” approach)

- thalamic nuclei present a promising target for RNS detection and stimulation for seizures with multifocal or generalized onsets.

Procedure – Depth Electrode implantation

Hippocampal depth

Trajectory (hippocampus cannulation) – see p. E15 >>

- prone in Leksell frame base (e.g. with special attachment to ROSA) or radiolucent Mayfield (but standard metal frame may have less wiggle than radiolucent frame) – make sure pins will not interfere with bur holes.

- guidance platform: Nexframe* or STarFix or ROSA**

*need O-arm registration (can be fiducialless)

**need O-arm registration (from frame pin ends, no need for bone fiducials)

Cannula options

- conventional DBS cannula – see p. Op360 >>

If using AlphaOmega microdrive system with the bengun and cannulas:

    - use “at target” cannula:

        - 1.81 mm outer diameter (each cannula slot on the bengun has diameter 1.8 mm)

        - 1.6 mm inner diameter

        - 192 mm length (set ROSA to “160 mm distance to target” - to accommodate the 30 mm height of non-XY CRW adapter).

    - if outer table-target length is > 120 mm, cannot use AlphaOmega microdrive system as will not be able to drive ROSA to “160 mm distance to target”, so rather use slotted cannula (see below).

    - use 44 cm depth leads!!! (30 cm lead will have trouble securing in DBS lead holder); except if outer table-target length is ≤ 80 mm may use 30 cm RNS depth lead and “10 mm above target” cannula – when microdrive is dialed all the way down, the distance between cannula distal tip and the first lead snap is about 29 cm (be prepared to use only one clip on DBS lead holder and clip the lead at stylet plastic handle by pulling to a side) and the distance between the proximal tip of cannula and the first lead snap is about 8 cm (so just enough to pull the cannula up and see the exposed lead at the outer table); otherwise, use 44 cm depth lead.

- slotted cannula is needed for STarFix or ROSA (unless using electrode holder when pulling regular non-slotted cannula up):

Two-piece slotted cannula (2SC-190X, Ad-Tech Medical Instrument Corp):

    - 1011  1200 USD (re-usable).

    - 190 mm length – set ROSA to “190 mm distance to target”

    - 2.11 outer diameter.

Incisions

for bilateral leads (keeping in mind scalp blood supply and the battery replacement surgeries):

A option:

B option (device on the right):

B option (device on the left):

Implantation with ROSA

  - O-arm registration.

  - 3.2 mm drill hole.

  - put a mark with permanent marker on a lead in two places:

- target to outer skull table

- length of cannula (i.e. top of ROSA platform) – not needed if using DBS lead holder

Implantation with Nexframe:

  - Acra-Cut perforator (undercut laterally using M8 drill bit).

  - Medtronic Nexframe needs double Silastic on lateral side (may even need offset bengun).

  - if using Nexframe: “at target” DBS cannula with Neuropace electrode is placed into the center track to the full depth to the target.

- do CT and look for contact position (10 mm lead with 4 contacts spans 32 mm) – may need to pull lead back for the best coverage.

Lead anchoring options

- Medtronic lead burhole anchoring device (Stimloc) – goes over standard bur hole 14 mm (old Medtronic leads 3389, 3387 have the same diameter as RNS depth leads – 1.27 mm; newer Medtronic SenSight DBS leads have diameter 1.36 mm and new type of burhole cover).

- Neuropace lead burhole anchoring device (Plunge) – goes over standard bur hole 14 mm.

- Dogbone titanium plate (advantage – skull twist hole can be very small; e.g. use 3.2 mm drill bit); to protect lead, slide a small piece of shipping tube over the lead where dogbone will sit (Neuropace formally recommends against using dogbone as lead fracture has been reported).

Implantation of 2 leads

- if implanting 2 leads in parallel (e.g. second lead in parahippocampal gyrus), aim to keep distance > 10 mm between leads so can stimulate in bipolar mode between contacts in separate leads.

- inner diameter of Neuropace bur hole cover is 10 mm (if using one bur hole for both leads, plan entry points ≤ 10 mm apart).

Procedure - Device Implantation

- two antibiotics: cefazolin, 2 g IV and vancomycin 1 g IV

- plan incision so generator is not under incision; may use pericranial flap of opposite direction.

- if incision is curved, plan generator implant so leads face away from incision; plan for future battery replacements – will need to open only scar at the flat end of device, curving a little bit on the side (opposite where leads exit).

Craniectomy

N.B. partial thickness craniectomy is not recommended – time consuming, bone grows and pushes device out.

- consider pericranial flap.

- use template (e.g. ferule) to mark bone.

- location – anywhere where shape conforms to skull curvature, e.g. parietal area.

- one or two bur holes (some place it where ferule tab is going to be).

- discard bone flap.

- some place dural tack-ups along perimeter.

Leads

- tunnel leads so the coil on the side (not above) the device, at the base of flap.

- electrodes are connected to Neuropace ports, e.g. the left hippocampal lead into port # 1, the left parahippocampal lead - into port # 2.

- normal lead impedance range 250-3500 Ohm.

Case report of implanting in infraclavicular position

A Novel Approach for Responsive Neural Stimulator Implantation With Infraclavicular Placement of the Internal Pulse Generator. Lucas R Philipp, Robert E Gross, MD, PhD Operative Neurosurgery, opy025, https://doi.org/10.1093/ons/opy025 Published: 14 March 2018

- band pass detection rates increased by 50%, while line length detection rates decreased by 50%.

- number of detections decreased from 1046 to 846, with a resultant decrease in stimulations.

- although there was some compromise of function due to the elevated noise floor, more than 2 yr following the procedure the patient remains free of seizures and infection.

Case report of implanting inside prosthetic skull implant

First In-Human Experience With Complete Integration of Neuromodulation Device Within a Customized Cranial Implant. Chad R Gordon, DO Gabriel F Santiago, MD Judy Huang, MD Gregory K Bergey, MD Shuya Liu, MS Mehran Armand, PhD Henry Brem, MD William S Anderson, PhD, MD. Operative Neurosurgery, Volume 15, Issue 1, 1 July 2018, Pages 39–45

Procedure - Battery Replacement

- enough to expose only very bottom (flat end) of generator → undo 2 screws (tab and connector cover), remove strain relief plate from leads, and reflect leads.

- use Bovie set at 6 if need to dissect leads (Dr. P. Weber); do not use Bovie when new battery is in.

Battery charge level

- patient uploads daily report to PDMS – report contains battery voltage; when low voltage is detected, Neuropace informs neurosurgeon about the need to schedule IPG replacement.

|  | Model 320 | Model 300 |
| --- | --- | --- |
| ERI (elective replacement indicated) | 2.37 V | 2.75 V |
| EOS (end of service) | 2.30 V | 2.70 V |

Procedure - Explantation

- explanting a chronically implanted cortical strip lead may cause cortical damage.

Complications

Infection (3.7-4.1%; 1.9% leading to explantations)

Hemorrhages (2.7%)

No chronic stimulation side effects!

Postoperatively

MRI

https://www.neuropace.com/rns-system-mri-guidelines/

- device gives MRI artefacts.

- MRI may induce low levels of current through implanted leads (tingling sensation), vibration, heating.

- MRI scan may be safely performed on patients with the RNS® System only under the specific conditions:

- > 10 days after surgery.

- no fever > 37 degrees

- only 1.5 Tesla horizontal, closed-bore MRI systems with spatial field gradient ≤ 30 T/m (3,000 gauss/cm), gradient slew rate ≤ 200 T/m/s per axis.

- stimulator has to be programmed into “MRI mode” (not detecting and not stimulating) and battery must not be in EOS.

Note: neurostimulator uses more battery power in MRI Mode (neurostimulator can be in MRI Mode for up to approximately two days per year without affecting battery longevity)

- only with model RNS-320 (RNS-300M is MRI unsafe) or incomplete RNS system (i.e. no IPG is implanted but other components still present):

N.B. all Neuropace lead models are eligible for MRI.

- leads can be attached or unattached to the neurostimulator.

- unattached leads can be capped or uncapped.

- leads that are cut or broken (at any point along the length of the lead) are eligible for MRI.

- do not use a head or extremity transmit coil

- active scan ≤ 30 minutes per session; wait 30 minutes between sessions.

Contraindicated procedures

- Diathermy procedures (any treatment that uses high-frequency electromagnetic radiation, electric currents, or ultrasonic waves to produce heat in body tissues); from RNS manual:

“Patients absolutely CANNOT be treated with any type of shortwave, microwave, or therapeutic ultrasound diathermy device whether or not it is used to produce heat. These treatments should not be applied anywhere on the body.”

- Electrosurgery – conflicting info: some say not to use Bovie (use bipolar at > 2-3 cm away from device); others say it is OK to use Bovie (if close to device, it may reset the device – will need to reprogram). It is OK to use Bovie on old battery (if close to leads – use lowest settings on Bovie, e.g. 6).

- Electroconvulsive Therapy (ECT).

- Transcranial Magnetic Stimulation (TMS).

- Radiotherapy

Exposure to high levels of radiation may damage the RNS® System.

The effects of high radiation sources (such as cobalt 60 or gamma radiation) on the RNS® System have not been studied - no studies to determine safe levels of irradiation for their device nor are there any recommendations for safe dose exposures. Neuropace suggests that the patient scan the device daily after each radiation fraction. This information can be transmitted to Neuropace who can determine if software repairs/re-programming can restore any lost functionality (if the device is damaged, it will try to reset itself. If unable to do that, it shuts down to a stable non-functioning state). Neuropace does not recommend prophylactically removing this device prior to radiation therapy.

Analysis & Programming

impedances - see above >>

Most important setting – where are detection contacts are.

Follow up schedule

- first 4 weeks postop - program just for recording as during this time recordings fluctuate and tend to gradually stabilize from surgical trauma.

- turn stimulation on at 4 weeks.

- clinic visits q3months until stabilizes.

RNS device stores ECoG only at certain triggers (if no triggers, device stores ECoG at certain time every day - scheduled):

- patient swipes magnet (e.g. if feels aura) – device stores 60 sec of preceding ECoG and 30 sec of subsequent ECoG.

- amplifier saturation (high amplitude)*

- long episodes (spikes or fast activity run > set time, e.g. 10 sec)*

*device stores ECoG if triggered stimulation does not abort activity (i.e. device stores ECoG of stimulation failures)

RNS-320 stores 8 ECoGs (90 second duration each - 60 sec of ECoG before trigger and 30 sec of subsequent ECoG after trigger); then it starts overwriting the oldest one without warning – so teach patient to regularly upload data to laptop – from RNS manual:

“The patient must collect data from the Neurostimulator once a day and send data to the PDMS once a week”

Detections

Detection counts: counts of irregular epileptiform activity detected by the RNS (mostly composed of brief interictal epileptiform events but also include a small number of electrographic seizure onsets).

Long episode counts: counts of a specific type of detection trigger (lasting longer than a pre-specified time period) that often represent electrographic seizures – could be a true objective ECoG biomarker for seizure burden.

David Spencer et al. Electrographic events recorded by a responsive neurostimulator may provide an objective assessment of clinical seizure burden.

Anderson, Daria Nesterovich (10/2023). "Closed-loop stimulation in periods with less epileptiform activity drives improved epilepsy outcomes.". Brain (London, England : 1878) (0006-8950)

- long episode counts significantly correlated with monthly baseline seizure rate.

N.B. only 10% of RNS triggers are actual seizures! (thus, it is correct to say that “RNS detects electrographic activity” and not “electrographic seizures”)

- interictal discharges – potential biomarkers.

- RNS detection threshold is 10 µV.

- program to detect seizures within 1 second.

Stimulation

Device stimulates up to 5 times (per FDA approval) if redetection happens right away.

  - current: 1-3 mA

  - pulse width: 160 µs

  - frequency: 100-200 Hz

  - burst duration: 100-200 ms

- avoid charge density > 4 µC/cm2 ← risk of preictal effect and patients do worse in studies; i.e. if charge density approaches 3 µC/cm2 but not improvement in seizure control, no reason to push higher.

  - when doing iEEG mapping, experts use up to 20 µC/cm2

  - RNS system allows to stimulate up to 25 µC/cm2

  - cellular damage happens when > 50-60 µC/cm2

- most RNS stimulation events happen in response to interictal epileptiform activity, which occurs hundreds to thousands of times daily compared to the seizures themselves, which occur at far lower rates → ratio of stimulation in the interictal vs. ictal period is > 1000:1 for most patients (i.e. RNS has chronic “disease-modifying” effect that is traditionally attributed to DBS).

Detection and Responsive Stimulation (hippocampal epileptiform discharge):

Outcomes

Neuropsychological

- some expert concerns that hippocampal (longitudinal) depth electrode placement may affect verbal memory – do neuropsychological testing 2 mos after implantation (before turning stim on).

- Loring et al. reported that there were no cognitive declines with responsive mesial temporal lobe stimulation and that there were improvements in verbal memory (small in magnitude but statistically significant) vs. progressive memory decline in nonoperated patients who continue to be treated with AEDs.

Loring DW et al. Differential neuropsychological outcomes following targeted responsive neurostimulation for partial‐onset epilepsy. Epilepsia 2015;56:1836–1844.

- statistically significant improvements in:

    - Naming (BNT; p<0.001)

    - Verbal learning (AVLT; p=0.03)

    - Visual memory (BVMT-R total recall; p=0.03)

    - Executive function [D-KEFS design fluency (p<0.001), WAIS-III Block design (p<0.001); WAIS-III Information (p=0.009)]

Quality of Life

- at 2 years of RNS treatment, there were clinically meaningful improvements in overall quality of life (QoL) in 41% of subjects, with only 16% reporting declines.

- Quality of Life Improvements:

Heck CN et al., Epilepsia, 2014

- patients treated with RNS earlier in the course of their epilepsy exhibited significant improvements in multiple mood and quality-of-life measures that were not seen in patients treated later in the course of their disease, despite similar efficacy in seizure reduction

Loring DW, Jarosiewicz B, Meador KJ, et al. Mood and quality of life in patients treated with brain-responsive neurostimulation: the value of earlier intervention. Epilepsy Behav 2021;117:107868.

Seizures

N.B. results improve over time!

- after RNS implantation, for 2 months postop seizure frequency may increase.

Geller 2017. RNS for temporal epilepsy

- RNS uncovered that clinical seizures are only the tip of iceberg.

Studies

1. Feasibility Study

- n=65

- open‐label

- assessed safety and followed adults over a 3‐month preimplantation baseline and a 2‐year postimplantation treatment period.

2. Pivotal Study

Heck CN, et al. Epilepsia. 2014 Mar;55(3):432-41

- randomized, controlled, double-blinded pivotal trial.

- enrolled adults with medically intractable partial epilepsy and an average of ≥ 3 partial seizures per month.

- n=191

- 3‐month baseline was followed by a 1‐month postimplantation stabilization period and then by a 4‐month blinded period during which subjects were randomized 1:1 to receive active or sham stimulation.

- all subjects received active stimulation in a subsequent open‐label period to collect safety and efficacy data to 2 years postimplantation.

3. Long-term treatment (LTT) trial

Bergey GK, et al. Neurology. 2015 Feb 24;84(8):810-7

Morrell et al. American Epilepsy Society. December 2016. Houston, TX.

- once subjects completed either the Feasibility or Pivotal study, they could continue in a follow‐on Long‐term Treatment (LTT) study.

- the study was ongoing as of the data cutoff on November 1, 2014.

- n=256

- at 7 years:

72% median seizure reduction

66% responder rate

at least one seizure-free period lasting:

≥ 3 months - 39% of patients

≥ 6 months - 29% of patients

≥ 1 year - 16% of patients

- at 9 years (1895 patient years):

75% median seizure reduction (35% of patients with ≥ 90% seizure reduction).

N.B. post-trial data shows 82% seizure reduction at 3 yrs

Razavi B, Rao VR, Lin C, et al. Real-world experience with direct brain-responsive neurostimulation for focal onset seizures. Epilepsia 2020;61(8):1749–57.

73% responder rate (≥ 50% seizure reduction)

at least one seizure-free period lasting:

≥ 6 months - 28% of patients

≥ 1 year - 18% of patients

no chronic stimulation-related adverse effects, no adverse cognitive or neuropsychological effects

- responder rate:

By seizure type

CPS		27%

GTC		65%

By ictal onset

Hippocampus	74%

Neocortex	37%

- responder rate by seizure focus:

Geller et al., Epilepsia, 2017; Jobst et al., Epilepsia, 2017

- seizure free periods by onset zone:

Geller et al., Epilepsia, 2017; Jobst et al., Epilepsia, 2017

- effectiveness is similar for:

- all ages and regardless of where the seizures started.

- subjects with and without mesial temporal sclerosis (MTS), bilateral MTL onsets, prior resection, prior intracranial monitoring, and prior VNS*

*mechanisms by which VNS and RNS act are likely quite different, suggesting that a failure to respond to one does not predict response to the other.

4. Postmarketing studies (Post-approval study, Real-world study)

Did improvements in RNS therapy has accelerated in the field’s post-marketing experience? The answer appears to be yes:

N.B. year 6 efficacy (of long-term study) achieved at year 1, year 9 efficacy – at year 2, superior efficacy at year 3

SUDEP

- SUDEP rate decreased at least 3-fold (2-3 per 1000 patient-years), compared with the expected rate in the DRE population (9 per 1000 patient-years).

Devinsky et al., Epilepsia. 2018: 1-7

*TRE = Treatment Resistant Epilepsy

**RNS System data represents SUDEP rate per 1000 stimulation years.
1 Ryvlin P, Cucherat M, Rheims S; Lancet Neurol. 2011; 10:961–8.
2 Dasheiff, R.M., 1991. J Clin Neurophysiol 8, 216–222.
3 Devinsky O, Friedman D, et al. Epilepsia. 2018; 1-7.

Medical Treatment Monitoring

- typically, it takes many months to know whether a new AED will improve seizure control.

- electrographic data recorded by the RNS may be used to assess whether a new AED is likely to improve seizure frequency in as early as one month.

Clinical and electrocorticographic response to AED with RNS (correlations between device data and clinical seizure frequency)

Tara L. Skarpaas, Thomas K. Tcheng, Martha J. Morrell. Clinical and electrocorticographic response to antiepileptic drugs in patients treated with responsive stimulation. Epilepsy & Behavior 83 (2018) 192–200

- in patients with RNS, adding clobazam or levetiracetam gave higher clinical benefit (seizure reduction) than starting lacosamide or pregabalin – that correlated with change in interictal spike rate.

Seizure prediction

Clinical seizures cluster in relation to cycles of interictal epileptiform activity Maxime Baud, Thomas Tcheng, Vikram Rao

- epilepsy is a cyclical disorder.

    - 60% of individuals exhibit multidien seizure cycles, and 89% have a circadian cycle.

Leguia MG, Andrzejak RG, Rummel C, et al. Seizure cycles in focal epilepsy. JAMA Neurol 2021;78(4):454–63

- using chronic ECoG from RNS and clinical seizure diary reports, interictal epileptiform activity (IEA) cycles were analyzed for 16 patients - all patients were found to have circadian and multi-day (“multidien”) IEA cycles with periodicities ranging from 7 to 33 days.

- an emerging view suggests seizures are not random events, and it may be possible to forecast clinical seizure risk on the scale of days with data obtained from the RNS.

12 months of RNS data from 1 patient – significant phase locking (clinical seizures occur in upward peak of IEAs):

Spencer et al. report circadian or ultradian patterns in 98% of subjects; data from 1 patient:

Spencer, Epilepsia, 2016.

Bibliography for ch. “Epilepsy and Seizures” → follow this link

Viktor’s Notes℠ for the Neurosurgery Resident

Please visit website at www.NeurosurgeryResident.net